## MRI Abdomen WWO
MRI Abdomen W/WO (Revised 12/06/21)
Sequence | Slice Thickness | Skip | Instructions
Cor SS FSE T2 | 5 | 1
Cor SS FSE TS F.S. | 5 | 1
Ax SS FSE T2 | 5 | 1
AX SS FSE T2 F.S. | 5 | 1
DWI | 7 | 0 | B Values of 50 and 800
Ax In & Out Phase | 5 | 1 | Separate and send by echo
Ax 2D Fiesta | 5 | 1
Cor Lava Pre | 3.6
Ax Lava Mask | 3.6 |  | Run Mask as pre-contrast
Ax Lava Post | 3.6 |  | Wait 20-23s post injection to start scan
Ax Lava Post | 3.6 |  | 1.5 Minutes post injection
Ax Lava Post | 3.6 |  | 3 Minutes post injection
Ax Lava Post | 3.6 |  | 5 Minutes post injection
Cor Lava Post | 3.6 |  | 5.5 Minutes Post injection (run last)
Coronal thin section MRCP and MRCP radial on all abdomens. |  |  | No need to do thin cuts through area of interest. i.e., pancreas, or adrenals. Just regular scan.
Without Contrast Abdomens obtain Ax and Cor Lava sequences. |  |  | System should be set up to subtract Pre from Post contrasted images. If this is not an option you need to do the subtraction manually and send it to PACs.(12/06/21)
Hepatic Fat Fraction is the "av=" number.

## MRI Abdomen WO
MRI Abdomen W/O (1/29/2021)
Sequence | Slice Thickness | Skip | Instructions
Cor SS FSE T2 | 5 | 1
Cor SS FSE TS F.S. | 5 | 1
Ax SS FSE T2 | 5 | 1
DWI | 7 | 0 | B Values of 50 and 800
Sag SS FSE T2 | 5 | 1 | If ordered for Risk Stratification for Polycystic Kidney Disease. Usually by Dr. Lindsay Crawford.
AX SS FSE T2 F.S. | 5 | 1
Ax In & Out Phase | 5 | 1 | Separate and send by echo
Ax 2D Fiesta | 5 | 1
Cor Lava | 3.6
Ax Lava | 3.6
Coronal thin section MRCP and MRCP radial on all abdomens. |  |  | No need to do thin cuts through area of interest. i.e., pancreas, or adrenals. Just regular scan.
Without Contrast Abdomens obtain Ax and Cor Lava sequences.
Hepatic Fat Fraction is the "av=" number.

## MRA Abdomen
MRA Abdomen (Revised 1/21)
Sequence | Slice Thickness | Skip | Instructions
Cor 2D Fiesta Fat Sat | 6 | 1
Ax 2D Fiesta Fat Sat | 8 | 1
Ax Inhance 3D In Flow IR | 2 | 0.5
Post
Cor CEMRA | 1.6 | 0.5 | Isotropic Voxel.
3-D Reconstructions | 19 views, 9.5 degrees between. |  | Rotate L-R and A-P

## MR Abd Elastography
MRI Abdomen with Elastography (Revised 05/18)
Sequence | Slice Thickness | Skip | Instructions
You will need an expiration breath hold scout in addition to your regular inspiration scout. The expiration scout will need to be hyperventilated so that they will be able to hold their breath out for the required time.
Cor SS FSE T2 | 5 | 1
Cor SS FSE T2 F.S. | 5 | 1
Ax SS FSE T2 | 5 | 1
AX SS FSE T2 F.S. | 5 | 1
Ax In & Out Phase | 5 | 1 | Separate and send by echo
Ax 2D Fiesta | 5 | 1
Cor Lava Pre |  |  | Run even if study is w/o only
Ax Lava Mask | 3.6 |  | Run even if study is w/o only
Ax Lava Post | 3.6 |  | Wait 20-23s post injection to start scan
Ax Lava Post | 3.6 |  | 1.5 Minutes post injection
Ax Lava Post | 3.6 |  | 3 Minutes post injection
Ax Lava Post | 3.6 |  | 5 Minutes post injection
Cor Lava Post |  |  | 5.5 Minutes Post injection (run last)
Coronal thin section MRCP and MRCP radial on all abdomens. | Must have order for both Abdomen and MR Elastography if they want both.(7/23/19)
Without Contrast Abdomens obtain Ax and Cor Lava sequences.

## MRI Abd Elastography PACS Info
Type | Images | Description
PROSP | 24 | Ax MR Touch ASSET
MR_ELASTOGRAPHY | 3 (send to PACS) | Elastogram + Mask (Gray)
SSAVE | 3 | Elastogram + Mask (0-8K)
SSAVE | 3 | Elastogram + Mask (0-20K)
MR_ELASTOGRAPHY | 3 | Elastogram (Gray)
SSAVE | 3 | Elastogram (0-8K)
SSAVE | 3 | Elastogram (0-20K)
SSAVE | 24 (send to PACS) | Wave Images
SSAVE | 1 (Spline ROI on Elastogram + Mask (Gray) (send to PACS) | SCREENSAVE

## MRI Athletic Pubaligia
MRI Athletic Pubalgia
To be performed at east on 3T if at all possible | 2017-01-27
Sequence | FOV | Slice Thickness | Skip | Matrix | Nex
Cor STIR | 28 | 4 | 1 | 256/192 | 2
Cor T1 | 28 | 4 | 1 | 256/256 | 1-2
Ax T2 F.S. | 28 | 4 | 1 | 256/256 | 2-3
Use PA torso or body coil for the above large FOV
Ax PD F. S. | 20 | 3 | 0.5 | 256/192 | 2-3
Ax PD | 20 | 3 | 0.5 | 256/192 | 2-3
Sag T2 F. S. | 20 | 3 | 0.5 | 256/256 | 3
Use surface coil placed at midline superficial to pubic symphysis for small FOV sequences. Reference attached pictures for detailed scan set-up.

## MRI Dynamic Pelvis
Dynamic Pelvis
2021-01-29 | Parameters are what is used on the 3T at East.
Sequence | FOV | Slice Thickness | Skip | Matrix | Nex | TR | TE | BW | ETL | Flip Angle | Special Parameters
AX T2 FRFSE XL | 25 | 5 | 0 | 352x256 | 2 | 5-6000ish | 102 | 35.71 | 19 |  | NPW,TRF, Fast, ZIP 512, FR
SAG T2 FRFSE XL | 25 | 5 | 0 | 352x256 | 2 | 5-6000ish | 102 | 35.71 | 19 |  | NPW,TRF, Fast, ZIP 512, FR
COR T2 FRFSE XL | 25 | 5 | 0 | 352x256 | 2 | 5-6000ish | 102 | 35.71 | 19 |  | NPW,TRF, Fast, ZIP 512, FR
DWI |  | 7 | 0 | [7 empty] | B Values of 50 and 800
SAG Dynamic Fiesta Resting | 36 | 20 | 0 | 384x256 | 1 |  | Min Full | 125 |  | 35 | Seq, EDR, Fast, MPh, ZIP 512
SAG Dynamic Valsalva | 36 | 20 | 0 | 384x256 | 1 |  | Min Full | 125 |  | 35 | Seq, EDR, Fast, MPh, ZIP 512
SAG Dynamic Keegle | 36 | 20 | 0 | 384x256 | 1 |  | Min Full | 125 |  | 35 | Seq, EDR, Fast, MPh, ZIP 512
SAG Dynamic Evac. | 36 | 20 | 0 | 384x256 | 1 |  | Min Full | 125 |  | 35 | Seq, EDR, Fast, MPh, ZIP 512
Patient uses cleansing enema at home prior to exam.
Have patient empty bladder prior to exam.
Use 60ml piston tip irrigation syringes.
Fill rectum with 120-150ml of water soluble gel with pt lying on side.
Fill vagina with 50ml of water soluble gel with pt supine. Use a SEPARATE syringe.
Provide baby wipes for patient after the exam.

## MRI Enterography
MRE | 2021-12-06 |  | NPO 4-6 hours prior
Sequence | Slice Thickness | Skip | Instructions
Cor T2 | 5 | 1
Ax T2 | 5 | 1
Cor FIESTA | 5 | 1
Cor T2 SS | 5 | 1 | Thru abd/pelvis
AX T2 SS | 5 | 1 | Thru abd/pelvis
Ax FIESTA | 5 | 1
DWI | 7 | 0 | B Values of 50 and 800
Ax Lava Pre | 3.6
Cor Lava Pre | 3.6
Cor Lava Post | 3.6 |  | 45 second post injection
Ax Lava Post | 3.6 |  | 45 second post injection
Ax Lava Post | 3.6 |  | 3 minute delay
Cor Lava Post | 3.6 |  | 3 minute delay
Ax 3D FIESTA
If MRCP ordered perform Coronal thin section MRCP and MRCP radial. | 1 Btl Volumen every 20 min, One Reglan Tab with First Btl, Glucagon Injection just prior to starting scan |  | If 14 years old or younger, treat as peds case and consult Dr. Verville.
Note in PACS: “MRE. 10MG REGLAN GIVEN PO, 1350ML VOLUMEN GIVEN P.O., 1ML GLUCAGON GIVEN I.M.”<---(or IV if applicable)[1/17/2020]

## MRI Hip
MRI Hip (Unilateral Hip Pain) Revised 2/16
Sequence | FOV | Slice Thickness | Skip | Matrix | Nex
Cor STIR (Sacrum to Pelvis) | 35-40 | 3 | 1
Cor T1 (Symptomatic Hip) | 16-18 | 3 | 0.5
Cor PD F.S. | 16-20 | 3 | 0.5
Ax PD F.S. | 16-20 | 3 | 0.5
Sag PD F.S. | 16-20 | 3 | 0.5
If Arthrogram add;
Cor T1 F.S. | 16-18 | 3 | 0.5
Cor T2 (No fat sat) | 16-18 | 3 | 0.5
Ax T2 (No fat sat) | 16-18 | 3 | 0.5
Use scanning parameters currently in use for each facility.

## MRI Pelvis for Anal Fistula
MRI Pelvis for Anal Fistula | 2021-01-29
Sequence | Slice Thickness | Skip | FOV | Instructions
Cor STIR | 5 | 1 | 40
Ax T2 SS | 5 | 1 | 22-26
Sag T2 SS | 5 | 1 | 22-26
AX SS FSE T2 F.S. | 5 | 1 | 22-26
DWI | 7 | 0 |  | B Values of 50 and 800
Ax LAVA Pre | 3.6 | 0 | 22-26
Cor LAVA Pre | 3.6 | 0 | 22-26 | Wait 20-23s post injection to start scan
Ax Lava Post | 3.6 | 0 | 22-26 | 1.5 Minutes post injection
Cor Lava Post | 3.6 | 0 | 22-26 | 3 Minutes post injection
Sag Lava Post | 3.6 | 0 | 22-26 | 5 Minutes post injection
Ax Lava Post Delayed | 3.6 | 0 | 22-26
All imaging should cover distally past perineum. No need to angle with the anal canal.

## MRI Prostate
MRI Prostate 11/12/19
Sequence | Slice Thickness | FOV | Skip | Instructions
Sag T2 FRFSE | 3 | 22 | Interleaved | Include Seminal Vessicles approx. 30 slices
Cor T2 FRFSE | 3 | 22 | Interleaved | 2 acquisitions. Approx. 25 slices. Angle with Prostate
Ax T2 FRFSE | 3 | 22 | Interleaved | Include Urethra. Approx. 60 slices
AX CUBE | 0.7 | 24
DWI | 6 | 28 | 0 | B50, 500, 1000, 1400. Copy scan coverage only. DWI - Functool - wide open - next, Advance Settings, Custom 0.50, Done, Compute Func maps. Please only send Calculated ADC maps to PACs.
PLEASE NOTE: IT IS VERY IMPORTANT THAT THE ADC MAP IS CALCULATED AT 0.2 AND SENT!! THIS IS THE
MOST IMPORTANT PART OF THE PROSTATE PROTOCOL!!!!
Ax LAVA | 1.6 | 28 |  | Run and check mask. Then start contrast and scan at same time. Needs to have an overall scan time of 2 minutes with as many phases that will allow that and adequate coverage of the anatomy. This should give you 6-7 phases. Please only send subtracted images to PACs, except for 1st non-subtracted scan.
AX T1 POST | 3 | 22-25 |  | Now required by Pi-Rads. To cover entire pelvis. Can be fast sequence.
Approved readers are Dr. Sud and Dr. Sabourin
Diffusion: 1. Set low end of parameters to 1. 2. Set Sensitivity to 0.2 under custom settings.
Before verifying!! *Make sure PSA levels are in PACs(a least the last 3) *Have BX reports, any radiation seed placement notes, Prostate Questionnaires, and Office Notes in PACs. This information is difficult to collect. Please have this information before the patient arrives for exam if possible. These should be scheduled at least 7 days out.
CIO 3T 16 CHANNEL ANTERIOR BODY COIL
CIE 3T 8 CHANNEL BODY COIL

## MRI Rectal Cervical Cancer
Rectal & Cervical Cancer
2018-07-19
Sequence | FOV | Slice Thickness | Skip
Ax T2 FSE (No F.S) | 16-18 | 3 | 0
Cor T2 FSE (No F.S) | 16-18 | 3 | 0
Sag T2 FSE (No F.S) | 16-18 | 3 | 0
Ax T2 SS FSE (No F.S.) | 30-36 | 4 | 1
Cor T2 SS FSE (No F.S) | 30-36 | 4 | 1
Sag T2 SS FSE (No F.S) | 30-36 | 4 | 1
AX T2 SS FSE F.S. | 30-36 | 4 | 1
Ax T2 CUBE 3D (If available) | 22-28 | 1.4 | 0
Ax LAVA Pre | 22-28 | 3 | 0
Ax LAVA Post (immediate) | 22-28 | 3 | 0
Cor LAVA Post (90sec. delay)
Ax Lava Post (90sec. delay)
Cor LAVA Post (3min. delay)
Ax LAVA Post (3min. delay)
DWI B0, 400, 800 calculate ADC
Warm US jelly 60cc for low & 80-100 for high rectal area of interest. | If coming from Shauna Lorenzo; Carr, Sabourin, Kyle, and Sud to read.
Liquid diet the day before and day of exam plus 1 bottle of mag citrate the night before.
Axial sequences need to be oblique, so axials are orthogonal to the tumor.

## MRI Soft Tissue Pelvis
MRI Soft Tissue Pelvis (Revised 4/21/21) |  | Use this for all pelvis studies except for prostate, dynamic pelvis and bony pelvis
Sequence | Slice Thickness | Skip | Instructions
Cor SS FSE T2 | 5 | 1
Cor SS FSE T2 F.S. | 5 | 1
Ax T2 SS | 5 | 1
Sag T2 SS | 5 | 1
AX SS FSE T2 F.S. | 5 | 1
DWI | 7 | 0 | B Values of 50 and 800
Ax In & Out Phase | 5 | 1
Ax LAVA Pre | 3.6 | 0 | Run Mask as pre-contrast
Cor LAVA Pre | 3.6 | 0
Ax Lava Post | 3.6 | 0 | Wait 20-23s post injection to start scan
Cor Lava Post | 3.6 | 0 | 1.5 Minutes post injection
Sag Lava Post | 3.6 | 0 | 3 Minutes post injection
Ax Lava Post Delayed | 3.6 | 0 | 5 Minutes post injection
Coronal thin section MRCP and MRCP radial on all abdomens. |  |  | No need to angle to uterus.